Useful Information to Run Behavioral Chamber

- All important files can be found under Documents→Behavior_Chamber
- Chamber design:
  - Chose Chamber_Design folder
  - New layers for chamber can be cut using laser cutter in Dickinson lab (see Laser_Cutter_Instructions)
  - Chamber design found in Layers folder. Layer 1 is top layer and should be cut from 1/16” clear acrylic. Layer 2 is second layer and should be cut from 1/32” white delrin. Layer 3 is third layer and should be cut from yellow Shim-Stock 0.20”. Layer 4 is bottom layer and should be cut from 1/16” clear acrylic. Frame is base layer to attach removable chamber to frame and should be cut from 1/8” clear acrylic. PCB_insert_small is layer where PCB board can be placed in and should be cut from 1/16” clear acrylic. LED_side_small is for mounting the LEDs, and should be cut from 1/8” white acrylic.
  - PCB board design can be found in PCB folder. Folder 1 has design for large board (where board is a full layer) and Folder 2 has design for small board (where board is inserted into PCB insert layer). There are still three extra small PCB boards on shelves above desk.
  - Chamber is held together by 8/32 x 1/2” stainless thumbscrews with wing nuts underneath for tightness. Odor line is attached via 10/32 UNF.
- Controls:
- Make sure all extension cords on top of computer are switched on
  - Chamber can either be operated manually or automatically.
  - For manual control, turn on MFCs and set air flow to desired value (instructions on how to do this can be found in MFC manual in top drawer by desk). Valves can be controlled by running NI_MAX desktop shortcut, clicking on Data Neighborhood → NI-DAQmx Global Virtual Channel → port0 and hit Run. Control valves 0-3 from here. Turn HPA on from wall. Control electric shock input by switching on power supply and turning output on when you want shock applied.
  - For automatic control of MFCs run code found in Controls→MFC_Controls→alicat_com. For automatic control of valves run code found in Contols→Valve_Controls→valve_control
- To track fly movements:
  - In theory you should be able to use Ctrax but I was unable to get this working
  - Manually score videos on either decisions for an odor (entry into odor half compared to choice zone) or percentage occupancy (record position in chamber for each fly at X number of frames). Record data in excel file under Data folder and analyze data under Analysis folder
- To load flies, screw two screws into top most holes in chamber (attaching all layers except top layer) and slide top layer so that vents provide space for loading.
- To prepare odors mix with parafin oil, using convention that 4ml parafin = 1.6g (weigh by mass of parafin not volume).
- See Previous_Methods for past papers that are useful and my final presentation and final report for more information on motivation/experiments run so far/future experiments.